LABKITPRO   |   ISO/IEC 17025:2017 toolkit   |   Uncertainty Evaluation

Measurement Uncertainty Budget

GUM (JCGM 100:2008) · generic worked example · not your data

DEMO SAMPLE — what this file is
A generic sample budget illustrating structure, calculation and reporting depth. The paid full version computes from YOUR components, applies your laboratory context, supports Monte-Carlo (JCGM 101) and exports an editable Word report.

1  Measurement model and method

Y = Σ cᵢ · xᵢ   — the default linear model used for this example. The full version accepts a user-defined model and applies sensitivity coefficients accordingly. Evaluation follows GUM (JCGM 100:2008); the expanded uncertainty uses a coverage factor k derived from the effective degrees of freedom, or from Monte-Carlo simulation (JCGM 101:2008) on request.

2  Uncertainty components (sample)

| Component | Type | Distribution | uᵢ | cᵢ | uᵢ·cᵢ | νᵢ |
| --- | --- | --- | --- | --- | --- | --- |
| Repeatability | A | normal | 0.0200 | 1 | 0.02000 | 9 |
| Reference standard | B | normal | 0.0150 | 1 | 0.01500 | 100 |
| Calibration of instrument | B | uniform | 0.0058 | 1 | 0.00580 | 50 |
| Resolution | B | uniform | 0.0029 | 1 | 0.00290 | 50 |

3  Combined and expanded result (sample)

| Quantity | Value | Basis |
| --- | --- | --- |
| u_c (combined standard) | 0.0250 | Root-sum-square of uᵢ·cᵢ |
| ν_eff (effective dof) | ≈ 38 | Welch–Satterthwaite |
| k (95% coverage) | 2.03 | Student-t at ν_eff |
| U = k · u_c | 0.0508 | ≈ ±5.1 % of reading |

4  Contribution analysis (sample)

| Component | Share of u_c² | Reading |
| --- | --- | --- |
| Repeatability | 64% | Dominant — increase replicates to reduce |
| Reference standard | 36% | Check the calibration certificate and its uncertainty |
| Calibration of instrument | <1% | Minor at this magnitude |
| Resolution | <1% | Minor at this magnitude |

5  Report statement (sample wording)

“The expanded uncertainty of the measurand is U = 0.0508 (k = 2.03, approximately 95 % confidence). The dominant contribution is repeatability; the second contribution is the reference standard.” The full version generates a statement of this form from your own measurand, unit, components and coverage factor, with the traceability chain recorded.

6  Traceability and records

Traceability — each input traceable to the SI through an unbroken chain of calibrations.

Records kept — budget, component data, calibration certificates, reviewer sign-off.

Review cycle — budgets re-validated when methods, equipment or scope change.

7  What the full deliverable adds

Computed from YOUR components — any number, any distribution.

Welch–Satterthwaite ν_eff and t-based k — or Monte-Carlo (JCGM 101) on request.

Context-linked report — aligned to your body, industry and assessment stage.

Optimisation pass — upload an existing report and see where U can be reduced.

Exportable, watermark-free Word (.docx) — ready to attach to your scope.

This is a generic demonstration sample produced by LabKitPro. It shows structure, wording and evidence logic only — none of your personalization inputs (country, accreditation body, industry, stage, uploaded documents) are applied. The full, tailored and watermark-free deliverable is generated from your context at labkitpro.com.